Modbus settings

| Baud | 19200 |
| --- | --- |
| Parity | None |
| Data Bits | 8 |
| Stop Bits | 2 |
| Interface | RS-485 3-wire |

Please be sure to only use the events recorded on the chamber.  E.g., if one chamber has RH, that will be the 2nd input process value, but if one does not have RH, then leave the 2nd input process alone.  Further, these values are to be read.

Here is an abbreviated list:

| English Name | Parameter | Address | Word Size | Representation | Enum Values | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| Temperature High Alarm Set Point | Input 1 High Alarm Set Point | 1306 | 2 | eREPRESENTATION_FLOAT |  |  |
| Temperature Low Alarm Set Point | Input 1 Low Alarm Set Point | 1308 | 2 | eREPRESENTATION_FLOAT |  |  |
| Backup Temperature High Alarm Set Point | Limit 2 High Alarm Set Point | 1700 | 2 | eREPRESENTATION_FLOAT |  |  |
| Backup Temperature Low Alarm Set Point | Limit 2 Low Alarm Set Point | 1702 | 2 | eREPRESENTATION_FLOAT |  |  |
| RH High Alarm Set Point | Input 2 High Alarm Set Point | 1312 | 2 | eREPRESENTATION_FLOAT |  | If needed on chamber |
| RH Low Alarm Set Point | Input 2 Low Alarm Set Point | 1314 | 2 | eREPRESENTATION_FLOAT |  | If needed on chamber |
| Temperature Alarm State | Input 1 Alarm State | 1339 | 1 | eREPRESENTATION_WIDE_ENUM | 88d = START, 61d = NONE, 12d = BLOCKED, 8d = ALARM LOW, 7d = ALARM HIGH, 28d = ERROR |  |
| RH Alarm State | Input 2 Alarm State | 1340 | 1 | eREPRESENTATION_WIDE_ENUM | 88d = START, 61d = NONE, 12d = BLOCKED, 8d = ALARM LOW, 7d = ALARM HIGH, 28d = ERROR |  |
| Backup Temperature Alarm State | Limit 2 Alarm State | 1704 | 1 | eREPRESENTATION_WIDE_ENUM | 88d = START, 61d = NONE, 12d = BLOCKED, 8d = ALARM LOW, 7d = ALARM HIGH, 28d = ERROR |  |
| Light Alarm State | Light Lifetime Alarm Status | 1345 | 1 | eALARM_2_MEMBER_LIGHT_LIFE_ALARM_STATUS | 63d=ON,62d=OFF,108=Onw/ACK |  |